## Slide 1: Crossover Distortion Hands-on Experiment TIPL 4101-L TI Precision Labs – ADCs
by Art Kay and Dale Li
Notes: Hello and welcome to the TI Precision labs hands-on experiment on crossover distortion. In this experiment we will compare the performance of two different rail-to-rail amplifiers. One amplifier, the OPA316 has input crossover distortion, and the other amplifier, the OPA320 does not have crossover distortion. In this experiment we will measure the degradation of THD caused by crossover distortion.

## Slide 2: Required/Recommended Equipment
Calculation
  Simple calculation using OPA320 and OPA316 Data Sheet
Simulation
  No simulation in this experiment.
Measurement
  PLABS-SAR-EVM-PDK
  http://www.ti.com/tool/plabs-sar-evm-pdk
  Download EVM software and purchase EVM
Notes: For this lab we will use the OPA320 and OPA316 data sheet to show if the device has crossover distortion, and the voltage level that the distortion occurs at. This experiment will not use simulation as the SPICE models do not model THD and crossover distortion. Finally, we will measure the two amplifiers. To make this measurement you will need the PLABS-SAR-EVM-PDK, and its associated software. The EVM can be purchased and the software can be downloaded using the provided link.

## Slide 3: Op Amp with and without input Crossover distortion
PARAMETER: OPA316 – Has Crossover |  | MIN | TYP | MAX | UNIT
VCM | Common Mode Voltage | (V-)-0.2 |  | (V+)+0.2 | V
CMRR | (V-)-0.2 < Vcm < (V+)-1.4V | 76 | 90 |  | dB
 | (V-)-0.2 < Vcm < (V+)+0.2V | 65 | 80 |  | 
PARAMETER: OPA320 – No Crossover |  | MIN | TYP | MAX | UNIT
VCM | Common Mode Voltage | (V-)-0.1 |  | (V+)+0.1 | V
CMRR | (V-)-0.1 < Vcm < (V+)+0.1V | 100 | 114 |  | dB
Much better CMRR 1.4V below positive rail. For 5.2V supply, crossover happens at 3.8V (5.2V – 1.4V = 3.8V)
Much better CMRR across entire common mode range.
Notes: The schematics at the left hand side show the circuits we will evaluate. The only difference between the two circuits is the op amp being used. The right hand side shows small excepts from the amplifiers data sheets. Specifically, we are looking at the common mode range and the common mode rejection, or CMRR specification. The common mode range for both amplifiers goes beyond the power supply rails. For the OPA316, the common mode range is 0.2V below the negative rail and 0.2V above the positive rail. In this example the supply rails are ground and 5.2V so the common mode range for the OPA316 is -0.2V to 5.4V. The OPA320 common mode range can be calculated similarly to be -0.1V to 5.3V.
Now let’s look at the common mode rejection specification. Although both devices are rail to rail the common mode rejection specification for the OPA316 is broken into two ranges. The first range shown is limited to 1.4V below the positive rail. Note that this restricted range has 10dB better performance than the full rail-to-rail range. On the other hand, the OPA320 only only lists one common mode rejection specification and the specified rejection is quite good, that is it is greater than 100dB.
The point here is that the OPA316 CMRR is better in the restricted range because it avoids the crossover region. The upper common mode limit here of 1.4V below the positive supply is about where the crossover distortion will happen. So in this case crossover happens 3.8V which is calculated by subtracting 1.4V from the 5.2V supply. The OPA320 on the other hand has good performance across the entire common mode range, so you can assume it is a zero input crossover distortion device. Also, note that most zero crossover amplifiers will highlight this feature on the data sheet front page, whereas rail-to-rail devices with crossover distortion may not emphasize this limitation.

## Slide 4: Connect the hardware
Notes: Now that we’ve done the simple calculation let’s set up the hardware and measure the performance. First set the jumpers on the PSI and the PLABS hardware as shown. Next plug the PHI into the PLABS channel 2 connection, and install the OPA316 into the channel 2 socket. Connect the PSI to the PLABS board using the SMA cable as shown. Finally plug the USB cables into your computer. When connecting the USB cables to the computer an LED on the PHI should blink and an LED on the PSI will illuminate continuously. Let’s pause for a moment and set the jumpers.

## Slide 5: Start & Setup the PLABS-SAR EVM Software
3. Set the sampling rate to 500ksps
1. Select “Plabs-SAR-EVM” from “start>All Programs”
4. Press here to expand PSI controls
2. If the EVM is operating correctly, “HW Connected” will be displayed.
Notes: Now let’s start the software by selecting the “Plabs-SAR-EVM” icon from the “Start > All Programs” menu. [click] Once the software is running you should notice the green “HW connected” message at the bottom of the software. [click] Next, change the sampling rate to 500ksps. [click] Finally press here to expand the controls for the PSI hardware.

## Slide 6: Power-on LEDs illuminate
8. Power on LED illuminate.
5. Blinking LEDs
6. Solid LED
7. Solid LED
Notes: When power is first connected to the PHI, you will see two blinking LEDs and one solid LED. [click] The PSI will have on solid LED. [click] Once the software boots up, you should notice the power LEDs on the PLABS board illuminate. [

## Slide 7: Setup the PSI
9. Enter the amplitude, and DC offset
10. Press Update.
11. Click to Turn on output.
Notes: Now let’s set up the PSI. First enter the peak to peak amplitude and DC offset required. Let’s start with 2Vpp and a 1.5V offset for this experiment. [click] Next press “Update” to save the changes. [click] Finally, click on this button to enable the PSI output. Note that the output is enabled when it is the teal color as shown, and it is disabled when it is red. Also notice that the graphical display indicates the output is about 0.5V to 2.5V as expected.

## Slide 8: Capture the waveform and zoom in.
12. Press Capture
13. This is a sine wave output, but the time scale doesn’t let you see the wave. Right click an drag to zoom in.
Notes: Press capture to read the ADC. [click] notice that the entire graph is filled with points. This is because the time scale isn’t adjusted properly and we need to zoom in to see a sine wave. Right click and drag to zoom in.

## Slide 9: Capture the waveform and zoom in.
14. Select “Fit Code to range” to show the full scale range.
Notes: Notice that the waveform fills the entire graph. This is because the graph automatically rescails is in the “Auto Mode”. Change to the “Fit code to range” mode, to see the waveform on the full scale range.

## Slide 10: Capture the waveform and zoom in.
15. Now you can see that the full scale range is displayed (5V) in this example.
17. Change to “Spectral Analysis”
16. Crossover happens at 3.8V for OPA316
Notes: Now you can see the waveform on the full scale range which is 5V in this example. Also, notice that the waveform range is 0.5V to 2.5V which matches what we entered using the PSI controls. [click] This display is useful as it helps us see that the signal is nowhere near the crossover region of 3.8V. [click] Now change the “page” to “Spectral Analysis”.

## Slide 11: Frequency Domain Results
19. Frequency domain display
20. Press “Mark Harmonics” to zoom in on harmonics.
18. Important AC measured results. Note that better SNR could be achieved with full scale signal.
Notes: Notice that under the spectral analysis page we can see important AC measured results like SNR and THD. Also, as a side note, the SNR for this experiment could be improved by applying a full scale signal. In this experiment, we are intentionally using a 2Vpp signal so that we can see the effects of crossover distortion. For best SNR we should really use a 5V signal. [click] Note that the frequency domain display shows the fundamental, noise floor, and harmonics. [click] To zoom in on the harmonics, click on “Mark Harmonics”.

## Slide 12: Mark Harmonics
21. Zoom in on harmonics marked H2 …H10
22. Amplitude of harmonics in table form.
23. Minimize or maximize PSI controls as needed.
Notes: Here we can see the zoomed in view of the harmonics. Notice that each of the harmonics is labeled H2 to H10. The display shows the nine harmonics that are used in the THD calculation. [click] Also notice that the harmonics are listed in a table format as well. [click] Click here to minimize or maximize the PSI controls as needed.

## Slide 13: Enter Vin, Vcm to compare OPA316 and OPA320
fin = 2kHz, fsamp = 500kHz Crossover region at 3.8V on OPA316 |  |  |  |  | 
PSI Signal Settings |  | Calculated PSI Min and Max Output |  | OPA316 Expected | 
Vin (Vpp) | Vcm (V) | Vmin (V) | Vmax (V) | SNR (dB) | THD (dB)
2 | 1.5 | 0.5 | 2.5 | 85.3 | -104.4
2 | 2 | 1 | 3 | 85.3 | -102.9
2 | 3 | 2 | 4 | 85.2 | -98.8
2 | 3.2 | 2.2 | 4.2 | 85.2 | -83.8
2 | 3.5 | 2.5 | 4.5 | 85.1 | -76.1
2 | 3.8 | 2.8 | 4.8 | 85.0 | -77.2
Notes: Now that we have gone through the hardware setup, we can fill out a table that compares AC performance of the OPA316 and OPA320 versus common mode voltage. To simplify collecting data, you can click on the number in the GUI cell and hit press control-C to copy the number. Then click in the table and press control-V to past the number. Also, for each row in the table, you will have to enter the amplitude from the Vin column, and the dc offset from the Vcm column. For example, for the first row we enter 2Vpp into the amplitude control, and 1.5V into the DC offset control.

## Slide 14: Measured vs Expected Results
fin = 2kHz, fsamp = 500kHz Crossover region at 3.8V on OPA316 |  |  |  |  |  |  |  |  |  |  | 
PSI Signal Settings |  | Calculated PSI Min and Max Output |  | OPA316 Expected |  | OPA316 Measured |  | OPA320 Expected |  | OPA320 Measured | 
Vin (Vpp) | Vcm (V) | Vmin (V) | Vmax (V) | SNR (dB) | THD (dB) | SNR (dB) | THD (dB) | SNR (dB) | THD (dB) | SNR (dB) | THD (dB)
2 | 1.5 | 0.5 | 2.5 | 85.3 | -104.4 |  |  | 85.7 | -109.2 |  | 
2 | 2 | 1 | 3 | 85.3 | -102.9 |  |  | 85.6 | -105.9 |  | 
2 | 3 | 2 | 4 | 85.2 | -98.8 |  |  | 85.7 | -106.1 |  | 
2 | 3.2 | 2.2 | 4.2 | 85.2 | -83.8 |  |  | 85.6 | -108.3 |  | 
2 | 3.5 | 2.5 | 4.5 | 85.1 | -76.1 |  |  | 85.6 | -109.1 |  | 
2 | 3.8 | 2.8 | 4.8 | 85.0 | -77.2 |  |  | 85.7 | -106.4 |  | 
Your results should show the same trend as the expected result but the specific values will differ.
Notes: This table shows the expected results for the OPA316 and OPA320. Let’s look at the first row. The input is set to 2Vpp, or 1Vpk. The common mode or offset is set to 1.5V. The sine wave negative peak is at Vcm – Vpk, in this case it’s 1.5V – 1V which is equal to 0.5V. Similarly, the positive peak is 1.5V +1V which is 2.5V. The positive peak, denoted Vmax, is the number we need to pay attention to. When Vmax is below the crossover voltage we will not see crossover distortion and the AC performance should be good. When Vmax is above the crossover voltage we should see crossover distortion and the AC performance should be degraded. The OPA316 crossover region happens at 3.8V, so we should see degradation above this voltage. Sure enough, you can see that the OPA316 THD degrades significantly for Vmax of 4V and greater. Notice, however, that the OPA320 has good performance across all common mode voltages because it is a zero-crossover device.
Now you can run the same experiment for the OPA320 and the OPA316. Your results should have the same trend as the expected results, but the specific value may differ somewhat. For example, you may see a THD of 100dB for Vmax of 2.5V, and 67dB for Vmax of 4.8V. The point isn't to focus on the exact value, but to notice that THD dramatically degrades once you enter the crossover region. If you don’t get these results, make sure that you have adjusted the hardware and software correctly. A common error is to forget to adjust the sampling rate to 500ksps. This is needed as the OPA316 is not fast enough for the ADS8860 full sampling rate of 1Msps.

## Slide 15: Graph of results
OPA316 Crossover at 3.8V
Vmax = 2.5V
OPA316 Crossover
3.8V
OPA316
OPA320
Your results should show the same trend as the expected result but the specific values will differ.
Click here to access Excel Spreadsheet for data collection.
Notes: Graphing the results from the table shows that the OPA320 THD is relatively unaffected by changing maximum input signal amplitude, but the OPA316 performance degrades when the maximum signal passes the crossover voltage of 3.8V. The attached Excel spreadsheet can be accessed by clicking on the icon in the power point presentation. This spreadsheet contains the table from the last slide and the graph shown here. When collecting data for the OPA320, use the coupon card labeled “OPA320_Good filter1”. All other jumper settings and connections are the same for both amplifiers.

## Slide 16
Thanks for your time!
Notes: That concludes the hands-on experiment. I hope this was useful to you. Thanks for your time.